| ECR: Implementing an OMC Front-End |  |  | ## DCC No: E2200441-v1 |
| --- | --- | --- | --- |
|  |  |  | ## Date: 11/1/2022 |
| Requester: Daniel Sigg | Impacted Subsystem(s): |  | ISC |
| Description of Proposed Change(s): Move the converters for the DCPD readout chain to a separate front-end and computer. More details in T2200373. |  |  |  |
| Reason for Change(s): Our DAQ front-ends experience intermodulation beats from the converter internal XO clocks when multiple boards of the same type are installed. These XO clocks are standard commercial units that have at least 25 ppm frequency uncertainty. For some the difference frequency produces a beat note that is visible in the GW readout. |  |  |  |
| Estimated Cost per site: Computer $3700; Adnaco $1120; Dolphin interface $810; Upgrade older FE chassis: 2000$; Timing board on hand; Backplane on hand; Power board on hand; Low noise ADC on hand; AA chassis $4000. Total ~$12k per site. |  |  |  |
| Schedule Impact Estimate: Tue maintenance? |  |  |  |
| Nature of Change (check all that apply): Hardware Safety Correct Hardware Correct Documentation |  | Improve Hardware/Software Improve/Clarify Documentation Change Interface Change Requirement |  |
| Importance: Desirable for ease of use, maintenance, safety Desirable for improved performance, reliability Essential for performance, reliability Essential for function Essential for hardware safety |  | Urgency: No urgency Desirable by date/event: January 2022 Essential by date/event: Immediately (ASAP): |  |
| Impacted Hardware (select all that apply): Repair/Modify. List part & SNs: ______________________ Scrap & Replace. List part & SNs: Installed units? List IFO, part & SNs: _________________ Future units to be built |  | Impacted Documentation (list all dwgs, design reports, test reports, etc.): ADC/DAC Channel Usage: T1100472 ISC/SQZ Wiring Diagram: D2200024 SQZ-C1 Rack: D1600513 |  |
| Disposition of the proposed change(s): The disposition of this proposed engineering change request is to be completed by Systems Engineering and indicated in the “Notes and Changes” metadata field in the DCC entry for this ECR. The typical dispositions are as follows: - Additional Information Required: in which case the additional information requested is defined. The ECR requester then re-submits the ECR with the new information using the same DCC number for the ECR but with the next version number. - Rejected: in which case the reason(s) for the rejection are to be given - Approved - Approved with Caveat(s): in which case the caveat(s) are listed - TRB: the ECR is referred to an ad-hoc Technical Review Board for further evaluation and recommendation. It is the System Engineer’s (or designee’s) responsibility to organize the TRB. The System Engineer (or designee) then makes a technical decision based on the TRB’s recommendation. Links to the TRB’s documentation (charge, memos, final report, etc.) are to be added to the “Related Documents” field for this ECR. - CCB: a change request for approval of additional funds or schedule impact is to be submitted to the Configuration Control Board. Links to the CCB’s documentation (CR, etc.) are to be added to the “Related Documents” field for this ECR. Concurrence by Project Management: Acknowledgement/acceptance/approval of the disposition is to be indicated by the electronic “signature” feature in the DCC entry for this ECR, by one the following personnel: - Systems Scientist - Systems Engineer - Deputy Systems Engineer |  |  |  |