## Slide 1
Photograph lit by scattered main beam light during high-sensitivity operation
lit by scattered main beam (observation sensitivity)
September 2023 PEM update
LIGO-G2301888
Robert Schofield, Anamaria Effler, Sheila Dwyer, Jane Glanzer, Adrian Helmling-Cornell, Georgia Mansell, Filiberto Clara, Richard McCarthy, Benjamin Mannix

## Slide 2
1) O3-O4 vibration coupling reduction projects successful
2) Does vibration coupling tend to increase with main laser power?
3) PEM injections successfully completed before O4 run at LLO, and at LHO at 75W – redo’s after LHO power drop to 60W almost complete.
4) Low vibration coupling at LHO and LLO Filter Cavity End Stations.
5) New electronic grounding noise injection technique reveals grounding problems.
6) There is a non-zero TM bias that minimizes electronic ground noise coupling.
7) Arm Cavity Baffle identified as scattering source at LLO, hanging problems corrected.
8) HVAC shutdown, LHO, increases range by about 10Mpc, HVAC coupling likely due to: A) coupling at EX cryobaffle, B) PSL (jitter), and C) unknown source of vibration coupling in LVEA.

## Slide 3
1) O3-O4 coupling reduction projects
annular reaction mass with reflective ESD traces
ETMY test mass
A) HAM5-6 septum window removal greatly reduced coupling, LHO, LLO
June 2022
December 2022
The estimated noise from ambient vibration is now mainly upper limits. I pushed the injections until the accelerometers started to saturate.

## Slide 4
Photographed while at full sensitivity, lit by scattered infrared beam.
1) O3-O4 coupling reduction projects
annular reaction mass with reflective ESD traces
ETMY test mass
  B) ETMY, ITMY Cryobaffle scattering noise reduced by damping;
  ETMX baffle not yet damped, LHO, LLO
4.0
5.0
5.5
7.7
8.2
actually on last rib, Y direction
Before ITMY Cryobaffle Damping
After ITMY Cryobaffle Damping
DARM
DARM
Accelerometer showing shaker frequency sweep
Accelerometer showing shaker frequency sweep
actually on last rib, Y direction

## Slide 5
1) O3-O4 coupling reduction projects
C) ITMY replacement reduced PSL jitter noise by almost 10, LHO
Right after ITMY replacement
Before, May 2019

## Slide 6
Photographed while at full sensitivity, lit by scattered infrared beam.
2) Does vibration coupling increase with laser power?
ETMY test mass
A) Input jitter coupling up by roughly 5 between 46W and 60W;
Strain
Input jitter, Yaw,
Black: nominal Red: injection on H1:IMC-PZT_YAW_EXC
60W March 15, 2023
annular reaction mass with reflective ESD traces
46W right after ITMY replacement May 12, 2021

## Slide 7
Photographed while at full sensitivity, lit by scattered infrared beam.
2) Does vibration coupling increase with laser power?
annular reaction mass with reflective ESD traces
ETMY test mass
B) scattering noise increased by ~10 from the EX cryobaffles at LHO from 38W to 75W, also at LLO
38W, Sept. 2020 showing noise resonance of ETMX
cryo-baffle at 4.1 Hz (LHO alog 56857)
75W May 2023 repeat of shaking showing that noise is generated at about the same frequency and Q, but with much greater amplitude relative to the background.

## Slide 8
2) Does vibration coupling increase with laser power?
C) What happened when we decreased power from 75W to 60W? PSL jitter coupling dropped by roughly two.
75W
100-1000 Hz speaker injection
60W

## Slide 9
2) Does vibration coupling increase with laser power?
C) What happened when we decreased power from 75W to 60W? Scattering noise at EX dropped by three.
DARM showing nominal spectra (thin lines) and during shaker injections (thick lines) at 75W (orange) and 60W (blue)
Peak comparison for similar velocities suggests that the scattered light amplitude is, at 60W, about 1/3 of the level at 75W

## Slide 10
2) Does vibration coupling increase with laser power?
Any indication of extra light at ETMX?
38W 2019
75W 2023
I tried to normalize the photos because they were taken with a factor of two power difference. I noticed no difference except perhaps in this region; if not a photographic artifact, it suggests that there is relatively more light that is outside of the beam, consistent with greater scattering.

## Slide 11
Photographed while at full sensitivity, lit by scattered infrared beam.
annular reaction mass with reflective ESD traces
ETMY test mass
Comparison of ACB photodiodes suggests that there is a moderate fractional increase in the small angle scattering detected by these diodes between 60W and 75W
2) Does vibration coupling increase with laser power?

## Slide 12
Photographed while at full sensitivity, lit by scattered infrared beam.
2) Do we have to worry about vibration coupling increasing with arm power?
annular reaction mass with reflective ESD traces
D) Possible mechanism: greater surface distortion from greater heating of coating defects may increase scattering and arm asymetry
Normalized noise
Input power (W)
Blue: LHO PSL jitter noise
Red: LLO EX scattering noise

## Slide 13
Photographed while at full sensitivity, lit by scattered infrared beam.
3) PEM injections completed before O4 run at LLO, and LHO at 75W – redo's after LHO dropped to 60W almost complete
annular reaction mass with reflective ESD traces
ETMY test mass

## Slide 14
Photographed while at full sensitivity, lit by scattered infrared beam.
4) Low vibration coupling at LHO and LLO Filter Cavity End Stations
annular reaction mass with reflective ESD traces
ETMY test mass
FCES accelerometers during acoustic injection of 5 Hz comb
Coherence
DARM

## Slide 15
5) New electronic ground injection and monitoring techniques reveal grounding problems
Injections of 11 Hz comb at EX onto ISC Common chassis with return from cable tray. The 1/f^3 slope in DARM is consistent with the 1/f slope in the current clamp monitor on the grounding cable.
Current monitor on grounding cable
EX: H1:PEM-EX_ADC_0_09_OUT_DQ
EY: H1:PEM-EY_ADC_0_14_OUT_DQ
CS: H1:PEM-CS_ADC_5_26_2K_OUT_DQ
Ground current monitor channels
Inject onto chassis
Return from cable tray
To set the scale:
For my ground injection at 22 Hz, I get about 0.3mV/m sqrt(Hz) at the EFM,
and, about 12 mV/sqrt(Hz) of fluctuation of the building ground, determined from my measurement of 2 ohms between building “ground” and earth, and the current measured by the current clamp on a grounding cable,
and, about 1.53 mV/sqrt(Hz) of drive fluctuation, as indicated by H1:SUS-ETMX_L3_DRIVEALIGN_L_OUT_DQ

## Slide 16
5) New electronic ground injection and monitoring techniques reveal grounding problems
Blue: before, Red: after chassis grounding improvements at EX
DARM
Coherence between DARM and current clamp on building grounding cable

## Slide 17
Photographed while at full sensitivity, lit by scattered infrared beam.
6) There is a non-zero test mass bias that minimizes grounding noise coupling.
annular reaction mass with reflective ESD traces
ETMY test mass
Time series of amplitude of 66 Hz peak in DARM from ground injection at EX as ESD bias is swept
Bias Voltage
DARM 66 Hz peak
128 V
EX
Bias sweep from -440V to +430V
Minimum measured in January 2023 ~130 V
Minimum, Aug, 2023, ~150 V

## Slide 18
6) There is a non-zero TM bias that minimizes grounding noise coupling.

## Slide 19
6) There is a non-zero TM bias that minimizes grounding noise coupling.
Current bias settings: ITMX -428V, ITMY -428V
Proposed bias settings: ITMX 0V, ITMY 60V
Dark Blue: clamp on grounding cable, Light Blue: magnetometer attached to BSC3, Red: rack magnetometer
Minimizing lines from electronic ground injections also reduced other electronic lines

## Slide 20
6) There is a non-zero TM bias that minimizes grounding noise coupling.
Possible mechanism: test mass is charged and bias is set so that force on charge cancels force on bias-polarized molecules
AI Chassis
ESD
ESD drive signal
Shield connected at both ends makes a local chassis ground among a group of chassis
Shields connected at ebay chasis
TCS Ethercat
2 ohms; note that this is ~50 parallel ground spike resistors, and not a single grounding spike
Ethercat
+ +
+
++
Ground potential fluctuations as current to ground varies
-+
-+
ESD drive fluctuations as local ground fluctuates with respect to building ground
Anodized racks make high-resistance connections to chassis ground
Force on bias-polarized molecules ( -+ ) can be adjusted to cancel force on excess charge ( + )
Neutral Earth
Mains transformer connected to building ground
Building ground potential varies by location depending on the low resistances in the grounding system
Ring heater
Viewport with fringing fields

## Slide 21
Photographed while at full sensitivity, lit by scattered infrared beam.
7) Arm Cavity Baffle identified as scattering source at LLO and hanging problems corrected
annular reaction mass with reflective ESD traces
ETMY test mass
Movies
transmission monitor (TMS)
ETMY test mass
arm cavity baffle (ACB)
cryo-baffle
lit by scattered main beam (observation sensitivity)

## Slide 22
Photographed while at full sensitivity, lit by scattered infrared beam.
8) LHO HVAC shutdown increases range ~10Mpc
annular reaction mass with reflective ESD traces
ETMY test mass
HVAC coupling likely due to:
A) coupling at EX cryobaffle,
B) PSL (jitter), and
C) unknown source of vibration coupling in LVEA

## Slide 23
1) O3-O4 vibration coupling reduction projects successful
2) Does vibration coupling tend to increase with main laser power?
3) PEM injections successfully completed before O4 run at LLO, and at LHO at 75W – redo’s after LHO power drop to 60W almost complete.
4) Low vibration coupling at LHO and LLO Filter Cavity End Stations.
5) New electronic grounding noise injection technique reveals grounding problems.
6) There is a non-zero TM bias that minimizes electronic ground noise coupling.
7) Arm Cavity Baffle identified as scattering source at LLO, hanging problems corrected.
8) HVAC shutdown, LHO, increases range by about 10Mpc, HVAC coupling likely due to: A) coupling at EX cryobaffle, B) PSL (jitter), and C) unknown source of vibration coupling in LVEA:

## Slide 24
105 m
160 m
78 m
160 m
Line for data on next slide

## Slide 27
transmission monitor (TMS)
annular reaction mass with electrostatic drive traces (concentric lines)
ETMY test mass
arm cavity baffle (ACB)
p-cal periscope
cryo-baffle
lit by scattered main beam (observation sensitivity)
Photographed while at full sensitivity, lit by scattered infrared beam.
Manifold/Cryopump Baffle: “Cryo-baffle”